## Slide 1: Load
Notes: Sample is loaded into the P7+ loop in 4 seconds. Carrier runs through the P6 valve

## Slide 2: Dilute
Notes: The valves switch and the syringes dispense for 6 seconds. The sample is mixed with the internal standard in valve. The mixed sample moves to the P6 loop (which is filled in 6 seconds).

## Slide 3: Inject
Notes: The P6 valve switches and the sample is introduced to the ICP / ICPMS and the P7+ sample loop begins its rinse cycle. The S2 – S4 syringes refill.

## Slide 4: Rinse
Notes: Rinse: 10 second rinse:
The valve on the S400V switches to rinse and the S2 carrier becomes S1 rinse. The vacuum is on and the sample and vacuum lines are evacuated: For the first 4 seconds the S1 pushes the P7+ loop into the P6 loop. After 4 seconds the P7+ valve switches it is evacuated.
S1 continues to push and cleans out the P6 loop for the remaining 6 seconds.